Quantification of Glyphosate and Other Organophosphorus Compounds in Human Urine via Ion Chromatography Isotope Dilution Tandem Mass Spectrometry

Andre Schütze, Pilar Morales, Meghan Vidal, Antonia M. Calafat, Maria Ospina

Division of Laboratory Sciences, National Center for Environmental Health, Centers for Disease Control and Prevention, 4770 Buford Hwy, MS S103-2, Atlanta, GA 30341, USA

Corresponding author: Andre Schütze, aschuetze@cdc.gov, 770-488-7845

Disclaimer

The findings and conclusions in this report are those of the authors and do not necessarily represent the views of the CDC. The use of trade names is for identification purposes only and does not constitute endorsement by the U.S. Department of Health and Human Services or the CDC.

Supplemental Material

| Column phase | Problem |
| --- | --- |
| reversed phase | no or little retention, little retention with interferences, carry-over |
| weak-anion exchange | interferences, sensitivity, carry-over (if metal casing or frits) |
| strong anion exchange | strong retention, solvents not compatible with MS, interferences, carry-over |
| HILIC | interferences, sensitivity, stability, carry-over |
| Non-conventional phases | TiO2, too much retention HCl required for elution (not MS compatible) |
|  | Mixed-mode anion and cation exchange mechanism present, interferences, carry-over, sensitivity |

Solvents used include water, methanol, iso-propanol and acetonitrile in combination with various buffers/modifiers ranging between pH 3 to 9 with a molar range of 5 mM to 100 mM include ammonium formate, ammonium acetate, sodium phosphate, disodium phosphate. Phosphate buffers are not MS compatible and were intended for the transfer between clean-up column and analytical column. We concluded that only a PEEK based system will be able to run without carry-over including the autosampler.