Supplemental Information

The output from the detector after amplification as measured with a 1 MΩ input impedance oscilloscope can be written as:

Oopen beam= 0∞Rλ*Tpassbandλdλ ≅  λ1λ2Rλdλ

Equation 1

where Rλ is the system response including the source, detector, and pre-amp characteristics. With the sample in the beam:

Osample= 0∞Rλ*Tsampleλ*Tpassbandλdλ  ≅ λ1λ2Rλ*Tsampleλdλ

Equation 2

if we assume the response of the system Rλ is constant from  λ1 to λ2

Oopen beam= C

Equation 3

Osample=Cλ1λ2Tsampleλdλ

Equation 4

For each passband filter, there is a different λ1,  λ1 and C, but for clarity we are only considering one passband filter at the moment.

The metric chosen, M, when utilizing the non-dispersive Infrared spectrometer has been the output from the detector with the sample in place divided by the output from the detector without the sample in place.

M ≡OsampleOopen beam

Equation 5

M≅λ1λ2Tsampleλdλ

Equation 6

This metric is measured for each of the passband filters.

We can approximate Tsampleλ and Tpassbandλ by measuring them with the FT-IR. In this way, we can arrive at an approximation of M.

M ≡OsamplemVOopen beammV≅ M= C λoλfTsampleλ*TpassbandλC λoλfTpassbandλ

Equation 7

To obtain Tpassbandλ, each passband filter was placed in the FT-IR and its transmittance measured; the results are shown in Figure 1. An example of Tsampleλ*Tpassbandλ is illustrated in Figure 2.

Figure S1: Transmittance of the 5 optical passband filters. Note filter 5 is outside the FT-IR measurement range.

Figure S2: Transmittance of passband filter, a sample (Teflon filter) as measured by FT-IR and their product.