| Supplemental Table 1 - Chromatography and MS instrument acquisition settings for PRM |  |  |  |
| --- | --- | --- | --- |
| Sample Volume | 20 µL, full-loop injection | MS2 Acquisition | PRM, 3 s AGC cycle time, centroid |
| Stationary Phase | Thermo Acclaim PepMap C18 75 µm × 50 cm | MS2 Fragmentation | HCD |
| LC Solvent A | 100% H2O, 0.1% formic acid | MS2 Detection | Orbitrap |
| LC Solvent B | 80% ACN, 0.1% formic acid | MS2 m/z range | 110-1100 m/z |
| Gradient Ramp Duration Flow Rate | 2.5-8% B in 3 minutes 3-25% B in 44 minutes 25-40% B in 4 minutes 300 nL/min | Isolation Window | 1.6 m/z |
| Mass Spectrometer | Thermo Orbitrap Fusion Lumos | MS2 AGC Target | 5e4 |
| Spray Voltage | 2.1 kV | MS2 Maximum IT | 54 ms |
| In-Source CID | 0.0 eV | Normalized Collision Energy | 30 |
| MS1 Scan Range | 325-1500 m/z | MS2 resolution | 30,000 @ 200 m/z |
| MS1 Resolution | 120,000 @ 200 m/z |  |  |
| MS1 AGC Target | 4e5 |  |  |
| MS1 Maximum IT | 50 ms |  |  |

| Supplemental Table 2 - Skyline and MS Analysis Settings for PRM |  |  |  |
| --- | --- | --- | --- |
| Digestion | Enzyme: Trypsin (0 max missed cleavages) | Precursor charges | 2,3,4 |
| Background Proteome | UniProt Rabbit Database (accessed November 29, 2018; 21177 entries) | Ion charges | 1,2 |
| Enforce Peptide Uniqueness | None | Ion types | b, y |
| Retention Time Predictor | None | Product ion selection | From: ion 3 To: last ion-1 |
| Ion Mobility Predictor | None | Special ions | N-terminal to Proline |
| Filter | Min length: 6 Max length: 30 Exclude N-terminal AAs: 1 | Ion match tolerance | 0.5m/z |
| Exclude peptides containing | Met, His, NXT, NXS, RP/KP | Instrument | Min m/z: 50 Max m/z: 2000 |
| Auto-select matching peptides | Yes | Method match tolerance m/z | 0.055 |
| Structural modifications | Carbamidomethyl | Alignment | Manual |
| Prediction | Precursor mass: Monoisotopic Product ion mass: Monoisotopic | Normalization (MS Stats) | Equalize Medians |
|  |  | Summary Method (MS Stats) | TMP Method |