## Slide 1: Supplemental Table 1S. Accuracy of CRMLN member reference laboratories, manufacturers and end-user clinical laboratories participating in CDC CVD BSP for 2013-2023 by analyte.
Analyte | Year From | Year To | Component | Mean Bias, % | Min Annual Mean Bias, % | Max Annual Mean Bais, % | Mean SE of Bias | Min Annual Median Bias, % | Max Annual Median Bias, % | Min Annual Bias IQR, % | Max Annual Bias IQR, %
TC-AK | 2013 | 2023 | CRMLN | -0.1 | -0.4 | 0.3 | 0.1 | -0.4 | 0.2 | 0.7 | 1.2
TC-AK | 2013 | 2023 | MFR | 0.2 | -1.0 | 1.0 | 0.1 | -1.0 | 1.0 | 1.9 | 3.4
TC-AK | 2013 | 2023 | LSP | -0.9 | -2.0 | 0.3 | 0.1 | -2.1 | 0.1 | 2.5 | 3.7
TG | 2016 | 2023 | CRMLN | 0.1 | -0.5 | 0.5 | 0.3 | -0.8 | 0.6 | 1.1 | 2.6
TG | 2021 | 2023 | MFR | 0.3 | -1.1 | 1.7 | 0.2 | -0.8 | 1.5 | 2.0 | 3.9
TG | 2013 | 2023 | LSP | -1.3 | -3.1 | 0.7 | 0.8 | -3.2 | 0.6 | 3.7 | 5.9
HDLC | 2013 | 2023 | CRMLN | -0.3* | -0.4* | 0.0* | 0.1* | -0.6* | -0.1* | 0.7* | 1.3*
HDLC | 2013 | 2023 | MFR | 0.3 | -0.4 | 2.0 | 0.1 | -0.4 | 1.0 | 4.4 | 7.7
HDLC | 2013 | 2023 | LSP | -1.1 | -2.3 | 0.0 | 0.1 | -2.8 | 0.1 | 5.5 | 6.8
LDLC | 2013 | 2023 | CRMLN | 0.0 | -0.6 | 0.4 | 0.2 | -0.6 | 0.4 | 1.3 | 2.7
LDLC | 2013 | 2023 | MFR | -0.2 | -2.7 | 2.3 | 0.2 | -2.4 | 2.3 | 3.9 | 6.6
ApoA-I | 2013 | 2021 | LSP | 1.0 | -0.6 | 2.5 | 0.1 | -0.7 | 4.0 | 5.2 | 6.9
ApoB | 2013 | 2021 | LSP | -3.7 | -6.3 | -1.1 | 0.1 | -6.7 | 0.0 | 4.9 | 9.0
*Expressed as mg/dL

## Slide 2
Analyte | Year From | Year To | Component | Mean CV, % | Mean SE of CV | Max Annual Median CV, % | Min Annual Median CV, % | Min IQR | Max IQR
TC-AK | 2013 | 2023 | CRMLN | 0.4 | 0.0 | 0.2 | 0.3 | 0.2 | 0.4
TC-AK | 2013 | 2023 | MFR | 0.6 | 0.0 | 0.3 | 0.6 | 0.4 | 1.0
TC-AK | 2013 | 2023 | LSP | 0.7 | 0.0 | 0.4 | 0.6 | 0.6 | 0.8
TG | 2016 | 2023 | CRMLN | 0.5 | 0.0 | 0.4 | 0.8 | 0.2 | 0.7
TG | 2021 | 2023 | MFR | 0.7 | 0.0 | 0.6 | 0.8 | 0.4 | 0.6
TG | 2013 | 2023 | LSP | 0.8 | 0.0 | 0.5 | 0.8 | 0.7 | 1.2
HDLC | 2013 | 2023 | CRMLN | 0.8 | 0.0 | 0.5 | 0.8 | 0.4 | 1.0
HDLC | 2013 | 2023 | MFR | 0.8 | 0.0 | 0.3 | 0.8 | 0.8 | 1.5
HDLC | 2013 | 2023 | LSP | 0.9 | 0.0 | 0.4 | 1.0 | 1.2 | 1.8
LDLC | 2013 | 2023 | CRMLN | 0.7 | 0.0 | 0.6 | 0.8 | 0.4 | 0.6
LDLC | 2013 | 2023 | MFR | 0.5 | 0.0 | 0.2 | 0.6 | 0.5 | 0.9
ApoA-I | 2013 | 2021 | LSP | 1.1 | 0.0 | 0.6 | 1.0 | 0.9 | 1.5
ApoB | 2013 | 2021 | LSP | 1.1 | 0.0 | 0.6 | 0.9 | 1.0 | 1.4
Supplemental Table 2S. Precision of CRMLN member reference laboratories, manufacturers and end-user clinical laboratories participating in CDC CVD BSP for 2013-2023 by analyte.