## Data Extraction
| Study | Location | Laboratory level | Degree of manufacturer involvement | Reference standard | Details of index test | Population | Range of CD4# | No. of samples (n) | Mean/ Median (cells/µl) | Range | Bias (cells/µl)* | LOA‡ (cells/µl) | Maximum observed difference† | Sensitivity, cut-off | Specificity, cut-off | Misclassification up (CD4 cut off, cells/µl) | Misclassification down (CD4 cut off, cells/µl) | No. of replicates per donor, no. of donors | CD4 count of replicate sample | %CV index, ref
[8 empty] | or |  | (SD) |  |  |  | (cells/µl) | [5 empty] | (cells/µl)
Subgroup
*MEDIAN
1 | Balakrishnan 2004 | Chennai, India | Charitable foundation clinical/research lab | One author affiliated to manufacturer | FACSCount | Coulter Cytospheres | HIV pos | ≈0-1150, mean 350.1 +/- 265.3 | 122 |  |  | 10 | -132 to +152 | +230, -250 |  |  | 5.1% | 3.6%
[17 empty] | -200 | -200
[8 empty] | <500 | ≈ 96 |  |  | (?+)8 |  | +120, -170
[8 empty] | >500 | ≈26 |  |  | (?+)15 |  | +230, -250
2 | Carella 1995 | USA | Teaching hospital / research lab? | None stated | DP flow | Coulter Cytospheres | HIV pos and neg | ≈0-1800, mean 534 +/- 509 | 117 | [7 empty] | 2.4% | 5.3%
[17 empty] | -200 | -200
3 | Landay 1993 | USA and Uganda | University research hospitals | 2 authors affiliated to manufacturer | DP flow |  | HIV pos and neg | ≈0-1600, 639 +/- 539 | 382 | [7 empty] | 16% | 2% | ? | 200 | 10.8%, 2.5%
[17 empty] | -200 | -200 |  |  | 3.5%, 2.5%
[19 empty] | ? | 1200
4 | Gernow 1995 | Denmark |  |  | DP flow | Coulter Cytospheres | HIV pos | ≈ 20-1000, 358 +/- 276 | 23 | [7 empty] | 12% | 20%
[17 empty] | -200 | -200
|  | Ivory Coast |  |  | DP flow |  | HIV pos | ≈10-2000, 540 +/- 407 | 41 | [7 empty] | 0% | 30% | ? 2, 19 (7 HIV pos, 12 neg) | Not stated | 57.7%, 6.0%
[17 empty] | -200 | -200
|  | Both sites | Not clear | Cytospheres donated by manufacturer | DP flow |  | HIV pos and neg | ≈10-5000 | 123 | [7 empty] | 7% | 20%
[9 empty] | (64 HIV pos, 59 neg) | [7 empty] | -200 | -200
[17 empty] | 27% | 5%
[17 empty] | -100 | -100
5 | Karcher 2006 | Uganda | Local lab, field conditions | None stated | DP flow | Coulter Cytospheres | HIV pos | ≈0-1200, median 332 | 131, from 102 donors |  |  | -1.4, | -372 to +369 | +425, -450 |  |  | 16% | 20%
[12 empty] | (+28M) |  |  |  |  | -350 | -350
[8 empty] | 0-500 | 91 |  |  | 61 | -192 to +315 | +400, -192
[8 empty] | <200 | 44 |  |  | +50.5M |  | +125, -125
[8 empty] | 201-350 | 23 |  |  | +59.5M |  | +400, -192
[8 empty] | 351-500 | 24 |  |  | +21.5M |  | +220,--190
[8 empty] | 501-750 | 27 |  |  | -69M
[8 empty] | >750 | 13 |  |  | -250M
6 | Lutwama 2008 | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | 0-900, 242 +/- 136 | 1444 (≈497 donors) |  |  | -56.8 | -225.8 to +112.2 | +240, -580 |  |  | 7% | 41%
[17 empty] | -200 | -200
[17 empty] | 1% | 68%
[17 empty] | -350 | -350
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | <350 | 1156 |  |  | -35.2 | -164.9 to +94.6 | +400,-580
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | <200 |  |  |  | 15M |  | +200,-230
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | 200-350 |  |  |  | 71M |  | +400,-580
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | 350-500 |  |  |  | 111M |  | +240,-440
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Coulter Cytospheres | HIV pos | >500 |  |  |  | 210M
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Dynal Dynabeads | HIV pos | ≈0-900, 239 +/- 140 | 1671 |  |  | -8.8 | -158.4 to +140.9 | +400, -500 |  |  | 13% | 16%
[17 empty] | -200 | -200
[17 empty] | 6% | 31%
[17 empty] | -350 | -350
|  | Uganda | University laboratory | None stated | FACSCount or FACSCalibur | Dynal Dynabeads | HIV pos | <350 | 1337 |  |  | -0.4 | -126 to +125.2 | +400,-270
[8 empty] | <200 |  |  |  | 1M |  | +340, -200
[8 empty] | 200-350 |  |  |  | 16M |  | +400,-270
[8 empty] | 350-500 |  |  |  | 24M |  | +410,--310
[8 empty] | >500 |  |  |  | 90M
7 | Idigbe 2006 | Nigeria | Research laboratory | None stated | FACSCount | Dynal Dynabeads | HIV pos and neg | ≈10-1600 | 97 (46 HIV pos, 51 neg) | [7 empty] | 17% | 20%
[17 empty] | -200 | -200
8 | Balakrishnan 2006 | Chennai, India | Charitable foundation clinical/research lab | None stated | FACSCount | Guava Easy CD4 | HIV pos | ≈10-1050, 303 +/- 230 | 110 |  |  | 26 | -44 to +96 | +125, -44 |  |  | 5% | 0%
[17 empty] | -200 | -200
[8 empty] | <200 | 24 |  |  | 10 | -22 to +43 | +55,-20
[8 empty] | 201-350 | 23 |  |  | 17 | -33 to +67 | +70,-44
[8 empty] | 350-500 | 23 |  |  | 36 | -38 to +110 | 100
[8 empty] | >500 | 20 |  |  | 60 | -28 to +148
[8 empty] | <350 |  |  |  | 13 | -27 to +53 | +70,-44
[8 empty] | >350 |  |  |  | 47 | +37 to +131
9 | Spacek 2006 | United States | University laboratory | None stated | DP flow | Guava Easy CD4 | HIV pos | ≈0-1800, 697 +/- 438 | 77 |  |  | -9 | -163.3 to +143.3 | +212.3, -294.1 |  |  | 11% | 1%
[17 empty] | -200 | -200
[8 empty] | <500 | 29 |  |  | -0.3 | -62.4 to +61.8 | +70,-130 | [8 empty] | * this Kandatil misclass data discovered late - added 12.07.2011
|  | Uganda | University laboraotory |  | FACSCount | Guava Easy CD4 | HIV pos | ≈0-1900, 335 +/- 331 | 141 |  |  | 4.6 | -187.6 to +196.8 | +479.9, -438.5 |  |  | 10% | 2%
[17 empty] | -200 | -200
10 | Pattanapanyasat 2007 | Bangkok, Thailand | Teaching hospital laboratory and ministry of public health-cdc lab | Manufacturer of reference test listed in sponsors | TruCount SP flow | Guava Easy CD4 | HIV pos | ≈0-1250 | 250 |  |  | 13.1 | -117.9 to +144.1 | +250, -300 |  |  |  |  | Within-run | Streck CD-Check Plus CD4 Low | <5%
?10,?1
[8 empty] | ≤200 | 76 |  |  | 8.9 | -27.5 to +45.3 | +55, -50 |  |  |  |  | Between-run |  | <8%
?2,?5
[8 empty] | >200 | 174 |  |  | 17.3 | -149.4 to +184.1 | +250, -300
[5 empty] | FACSCount | Guava Easy CD4 | HIV pos | ≈0-1200 | 250 |  |  | 33.2 | -101.8 to +168.3 | +300, -270
[8 empty] | ≤200 | 76 |  |  | 16.3 | -20.6 to +53.1 | +65, -20
[8 empty] | >200 | 174 |  |  | 44.9 | -122.6 to +212.3 | +300, -270
11 | Pattanapanyasat 2008 | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | TruCount SP flow | Guava Easy CD4 | HIV pos | 1-1201, 313.50 +/- 260.06 | 150 |  |  | 6.3 | -125.3 to +137.9 | +200, -300 |  |  |  |  | Within-run |  | 5.6%,3.7% (TruCount), 2.9% (FACSCount)
5,10 (fresh)
and 3,5 (CDCheckPlus CD4 Low)
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | TruCount SP flow | Guava Easy CD4 | HIV pos | ≤200 | 49 |  |  | 12.3 | -23.1 to 47.6 | +55, -28 |  |  |  |  | Between-run |  | 9.0%, 6.6% (TruCount), 5.8% (FACSCount)
3,5 (stabilised)
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | FACSCount | Guava Easy CD4 | HIV pos | 1-1137, 295.73 +/- 246.08 | 150 |  |  | 24.1 | -107.4 to +155.5 | +225, -260
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | FACSCount | Guava Easy CD4 | HIV pos | ≤200 | 49 |  |  | 18.4 | -20.1 to 57.0 | +65, -18
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | TruCount SP flow | CyFlow green | HIV pos | 1-1201, 313.50 +/- 260.06 | 150 |  |  | -27.6 | -117.2 to 62.1 | +175, -230 |  |  |  |  | Within-run |  | 6.7%, 3.7% (TruCount), 2.9% (FACSCount)
5,10 (fresh)
and 3,5 (CDCheckPlus CD4 Low)
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | TruCount SP flow | CyFlow green | HIV pos | <200 | 49 |  |  | -5.4 | -28.9 to 18.1 | +12, -35 |  |  |  |  | Between-run |  | 9.2%, 6.6% (TC), 5.8% (FACSCount)
3,5 (stabilised)
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | FACSCount | CyFlow green | HIV pos | 1-1137, 295.73 +/- 246.08 | 150 |  |  | -9.8 | -81.9 to 62.3 | +180, -140
|  | Thailand | Teaching hospital laboratory | Manufacturer of reference test listed in sponsors | FACSCount | CyFlow green | HIV pos | <200 | 49 |  |  | 0.8 | -21.7 to 23.2 | +18, -28
12 | Karcher 2006 | Uganda | Local lab, field conditions | None stated | DP flow | Partec Cyflow | HIV pos | ≈0-1050, median 288 | 128 from 121 donors |  |  | 92 | -217 to +402 | +500, -230 |  |  | 29% | 7%
[12 empty] | (+68M) |  |  |  |  | -350 | -350
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | 0-500 | 92 |  |  | 103 | -149 to +355 | +500, -210
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | <200 | 44 |  |  | +70M |  | +250,-40
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | 201-350 | 29 |  |  | +100M |  | +450, -200
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | 351-500 | 19 |  |  | +92M |  | +380,-75
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | 501-750 | 28 |  |  | +32.5M
|  | Uganda | Local lab, field conditions | None stated | DP flow | CyFlow Counter CD4/SCC no-lyse no-wash | HIV pos | >750 | 8 |  |  | +178M
13 | Zijenah 2006 | Zimbabwe | University laboratories | None stated | FACSCount | CyFlow SL | HIV pos | 0-1200, median 371, IQR 240-513 | 150 |  |  | 1.21 | -40 to +42 | +55,-60 |  |  | 4%
-200
[8 empty] | 0-250 | 40 |  |  | 2.6 | -21 to 27 | +30,-23
[8 empty] | 251-500 | 69 |  |  | 0.96 | -38 to 40 | +35,-50
[8 empty] | 501-1200 | 41 |  |  | 0.29 | -55 to 55
14 | Pattanapanyasat 2005 'Evaluation of a new..' | Bangkok, Thailand | Teaching hospital laboratory | None stated | TruCount SP | CyFlow Green | HIV pos | ≈0-1550, 360.5 +/- SD 309.1 | 200 |  |  | -69.1 | -225.7 to +87.5 | +180, -400 |  |  |  |  | Within-run | CD-Check Plus CD4 Low | 2.8%
[6 empty] | (single parameter CD4) | [12 empty] | 10,1
|  |  | Teaching hospital laboratory | None stated | TruCount SP | CyFlow Green (single parameter CD4) | HIV pos | <250 | 92 |  |  | -21.8 | -71.1 to +27.5 | +20, -120 |  |  |  |  | Between-run | CD-Check Plus CD4 Low | 7.8%
15,1
|  |  | Teaching hospital laboratory | None stated | TruCount SP | CyFlow Green (single parameter CD4) | HIV pos | <200 | 81 |  |  | -17.3 | - 57.9 to +23.3 | +20,-70
|  |  | Teaching hospital laboratory | None stated | FACSCount | CyFlow Green (single parameter CD4) | HIV pos | ≈0-1400, 331.4 +/- SD 291.7 | 200 |  |  | -40 | -165.1 to +85.1 | +300, -350
|  |  | Teaching hospital laboratory | None stated | FACSCount | CyFlow Green (single parameter CD4) | HIV pos | <250 | 96 |  |  | -7.2 | -42.9 to 28.4 | +36, -84
|  |  | Teaching hospital laboratory | None stated | FACSCount | CyFlow Green (single parameter CD4) | HIV pos | <200 | 85 |  |  | -5.8 | -37.6 to +25.9 | +40, -80
15 | Schnizlein-Bick 2000 | USA | 5 NIAID certified FC laboratories | Some authors affiliated to manufacturer | DP flow | SP FC using TruCount beads | HIV pos | ≈one third in each stratum: <200, 200-550, >500 | 60 common samples shipped to 5 sites, n=411 |  |  | +7M | [6 empty] | Inter-laboratory: | Median 315 | 9%,16%
5,60
[8 empty] | <200 | 168 |  |  | +2M | [6 empty] | Inter-laboratory: | Median 130 | 11%, 28%
5,21
[8 empty] | 200-500 | 156 |  |  | +17M | [6 empty] | Inter-laboratory: | Median 336 | 8%,15%
5,24
[8 empty] | >500 | 87 |  |  | +12M | [6 empty] | Inter-laboratory: | Median 714 | 9%,10%
5,15
[8 empty] | Half <200, half ≥200 | 14 local donors at each of 5 sites, 8 replicates of each sample tested by index and ref test (n=560?) |  |  | +10M | [6 empty] | Intra-laboratory: |  | 7%,7%
8,70
[8 empty] | <200 | n=280? |  |  | 5 | [6 empty] | Intra-laboratory: |  | 8%,11%
8,35
[8 empty] | ≥200 | n=280? |  |  | 24 | [6 empty] | Intra-laboratory: |  | 6%,6%
8,35
16 | Reimann 2000 | USA | 5 clinical FC laboratories | Partially funded by manufacturer | DP flow | SP FC using Flow-Count fluorospheres | HIV pos | Local specimens: | ≥14 from each of 5 sites=71 | [9 empty] | Within-laboratory precision: | All (see range) | 5.6,5.8
[8 empty] | ≈40% <200, 60%>200 | [10 empty] | 8,71 | <200
[20 empty] | >200 | 7.4,8.9
5.2,5.1
[6 empty] | Two colour + flow count |  | Central specimens: ≈37% <200, 50% 200-500, 13% >500 | 67 |  |  | Median differences of between -47 and +7 seen across different laboratories | [6 empty] | Between-laboratory precision: | All (see range) | 10.8,18.2
[19 empty] | 5,67 | <200
[20 empty] | 200-500 | 12.0,19.5
[20 empty] | >500 | 9.4,16.7
12.1,18.8
[5 empty] | DP flow | tetraONE system | HIV pos | Local specimens: | ≥14 from each of 5 sites=71 | [9 empty] | Within-laboratory precision: | All (see range) | 4.8,5.8
[8 empty] | ≈40% <200, 60%>200 | [10 empty] | 8,71 | <200
[20 empty] | >200 | 4.6,8.9
4.9,5.1
[8 empty] | Central specimens: ≈35% <200, 48% 200-500, 17% >500 | 71 |  |  | Median differences of between -36 and +19 seen across different laboratories | [6 empty] | Between-laboratory precision: | All (see range) | 9.6,18.2
[19 empty] | 5,71 | <200
[20 empty] | 200-500 | 9.6,19.5
[20 empty] | >500 | 10.8,16.6
9.5,19.8
17 | Storie 2004 | UK | Teaching hospital laboratory? | None stated | TruCount SP | SP FC using Perfect Count beads | HIV pos | 21-3163 | 104 |  |  | 27 | -67 to 121
[8 empty] | 0-200 | 24 |  |  | 9 | -13 to 31
[8 empty] | 201-500 | 39 |  |  | 22 | -34 to 78
[8 empty] | 501-1000 | 22 |  |  | 37 | -23 to 97
[8 empty] | >1000 | 19 |  |  | 47 | -139 to 233
18 | Storie 2003 | UK | Teaching hospital laboratory | None stated | SP flow with TruCount beads | SP FC using flow rate calibration | HIV pos | ≈0-3000 | 113 |  |  | -8 | -123 to +108 | -230,+160
[8 empty] | 0-200 | 25 |  |  | 0 | -29 to +28 | -30, +30
[8 empty] | 201-500 | 40 |  |  | -7.4 | -62 to +48 | -100, +50
[8 empty] | 501-1000 | 27 |  |  | -4 | -90 to +82 | -120,+105
[8 empty] | >1000 | 21 |  |  | -22 | -260 to +216 | -230,+160
[5 empty] | SP flow with TruCount beads | DP PLG | HIV pos | ≈0-3000 | 113 |  |  | -28 | -154 to +99 | -330, +140
[8 empty] | 0-200 | 25 |  |  | -5.4 | -32 to +22 | -50,+20
[8 empty] | 201-500 | 40 |  |  | -17.4 | -60 to +25 | -80,+20
[8 empty] | 501-1000 | 27 |  |  | -28.8 | -125 to +68 | -160,+75
[8 empty] | >1000 | 21 |  |  | -73.6 | -320 to +173 | -330,+140
19 | Jani 2011 | Mozambique | Primary healthcare clinics | None stated | SP flow on FACSCalibur | Pima (results with capillary blood extracted here; study also looked at venous blood) | HIV positive | 19-1450 | 135 |  |  | -52.8 | -250.9 to +145.2 | -620, +100 |  |  | 0% (200) | 5.2% (200) | Within lab precision: 2,103 (index: capillary blood collected into EDTA then pipetted into 2 PIMA cartridges); 2,140 (ref) | Mean 373.6 (index), mean 427.7 (ref) | 10.7%, 7.5% | (Also looked at pima venous vs ref, pima venous vs pima capillary)
[8 empty] | <500 | 89 |  |  | -24.4 | -127.9 to +79.2 | no plot |  |  | 2.2% (350) | 14.8% (350)
[8 empty] | >500 | 46 |  |  | -107.9 | -386.7 to +170.9 | no plot
20 | Sukapirom 2011 | Bangkok Thailand | Teaching hospital laboratory | None stated | FACSCount | Pima (venous blood in EDTA) | HIV positive | 4-2041 (mean 500 +/- SD 298) | 203 |  |  | -54.2 | -190.9 to +82.5 | [5 empty] | Within lab precision: 2,203 (total, unclear how many <200) | 0-200 | 2.84% | (Also looked at % similarity)
[8 empty] | <350 | 62 |  |  | -16.6 | -88.4 to +55.3 | [5 empty] | 2,62 | 0-350 | 13.92%
[8 empty] | >350 | 141 |  |  | -70.7 | -216.5 to +75.0 | [5 empty] | 2141 | >350 | 12.79%
[5 empty] | DP FC (FACSCan) | Pima (venous blood in EDTA) |  | 8-2036 (mean 490 +/- SD 297) | 203 |  |  | -44.0 | -179.7 to +91.6
[8 empty] | <350 | 62 |  |  | -18.0 | -89.4 to +53.4
[8 empty] | >350 | 141 |  |  | -55.5 | -206.0 to +95.1
21 | Renault 2010 | Burkina Faso | Community laboratory | None stated | FACSCount | Guava Easy CD4 | HIV positive | 0-1100 | 98 |  |  |  |  | +350, -320 |  |  | 6.1% (350) | 9.4% (350)
[8 empty] | <200 | 39 |  |  | +45.5
[8 empty] | 200-350 | 27 |  |  | +25.1
[8 empty] | >350 | 32 |  |  | -45.3
22 | Sitoe 2011 | Maputo City, Mozambique | Not stated | None stated | FACSCount using venous blood | FACSCount using capillary blood | HIV positive adults | Range not stated. 25th to 75th percentile: 218.5-412.5 | 101 |  |  | +4.3 | -61.7 to +70.3 | [5 empty] | 3,9 |  | 3.6% (index), 3.9% (ref)
[8 empty] | <350 | 65 |  |  | +9.9 | -34.2 to +54.1
[8 empty] | 350-500 | 27 |  |  | -5.5 | -83.9 to +72.8
[8 empty] | 501-999 | 9 |  |  | -7.3 | -128.4 to +113.7
[8 empty] | >1000 | 0
[5 empty] | FACSCalibur using venous blood | FACSCalibur usng capillary blood | HIV positive adults | Range 21-782. 25th to 75th percentile: 205.0-391.5. | 101 |  |  | +18.4 | -60.2 to +96.9 | [5 empty] | 3,9 |  | 5.5% (index), 7.4% (ref) | Also reportd % bias
[8 empty] | <350 | 64 |  |  | +18.2 | -48.6 to +85.1
[8 empty] | 350-500 | 28 |  |  | +18.5 | -80.7 to +117.6
[8 empty] | 501-999 | 9 |  |  | +19 | -69.7 to +107.7
[8 empty] | >1000 | 0
[5 empty] | FACSCalibur | FACSCount | HIV positive adults and children | 21 to 3073 | 152 |  |  | +6.7 | -158.6 to +171.9 | These data are for 'bias between the two instruments', unclear which way the bias lies (ie which instrument used as index or ref)
[8 empty] | <350 | 65 |  |  | -13.1
[8 empty] | 350-500 | 28 |  |  | -9.1
[8 empty] | 501-999 | 29 |  |  | -0.7
[8 empty] | >=1000 | 30 |  |  | +69.2
23 | Bergeron 2012 | Overall | National Laboratory | None stated | Single Platform Flow cytometry | PointCare NOW CD4 System | HIV pos and HIV neg | Overall | 472 | 325* | 4-1613 | (+) 153 | 221* |  | 53%, 350: 39%, 200 | 94%, 350: 94%, 200
[8 empty] | <350 | 259 | 220* | 4-350 | (+) 180 | 204*
[8 empty] | >350 | 213 | 517* | 351-1613 | (+) 120 | 236*
|  | A. Mozambique | National Laboratory | None stated | FACSCalibur with TruCount beads | PointCare NOW CD4 System | HIV positive | Overall | 143 | 252* | 9-1291 | (+) 237 | (-)402.2-769.2 | 1100 | 49%, 350: 32%, 200 | 80%, 350: 85%, 200 |  |  | intra-run variability + inter-run variability
[8 empty] | <350 | 101 | 194* | 9-350 | (+) 226 | 252*
[8 empty] | >350 | 42 | 459* | 353-1291 | (+) 81 | 373*
|  | B.Belgium | National Laboratory | None stated | FACSCalibur with TruCount beads | PointCare NOW CD4 System | HIV positive | Overall | 114 | 317* | 4-918 | (+) 212 | (-)160.3-584.0 | 800 | 38%, 350: 26%, 200 | 96%, 350: 96%, 200 |  |  | intra-run variability
[8 empty] | <350 | 68 | 227* | 4-350 | (+)220 | 146*
[8 empty] | >350 | 46 | 529* | 356-918 | (+) 200 | 241*
|  | C.Canada | National Laboratory | None stated | FACSCalibur with FlowCount beads | PointCare NOW CD4 System | HIV positive | Overall | 89 | 430* | 22-1272 | (+) 150 | (-)175.4-474.7 | 500 | 50%, 350: 44%, 2200 | 100%, 350: 98%, 200 |  |  | inter-run variability
[8 empty] | <350 | 26 | 250* | 22-345 | (+) 181 | 163*
[8 empty] | >350 | 61 | 506* | 350-1272 | (+) 141 | 167*
|  | D. USA | National Laboratory | Not stated | FACSCalibur with TruCount beads | PointCare NOW CD4 System | HIV negative | Overall | 55 | 541* | 109-1613 | (-) 11 | (-)161.9-157.2 | 175 | 100%, 350: 100%, 200 | 97%, 350: 94%, 200
[8 empty] | <350 | 21 | 295* | 109-342 | (-) 47 | 61*
[8 empty] | >350 | 34 | 762* | 359-1613 | (+) 11 | 91*
|  | E. South Africa | National Laboratory | None stated | Epics-XL with FlowCount beads | PointCare NOW CD4 System | unknown HIV status | Overall | 71 | 301* | 29-1332 | (+) 126 | (-)144.0-395.8 | 425 | 36%, 350: 50%, 200 | 94%, 350: 100%, 200 |  |  | inter-run variability
[8 empty] | <350 | 41 | 221* | 29-350 | (+) 124 | 115*
[8 empty] | >350 | 30 | 469* | 356-1332 | (+) 128 | 166*
24 | Dieye 2011 | Senegal | Reference laboratory | None stated | FACSCount | Auto 40 system with Easy MoAb CD4 reagents | HIV positive and HIV negative | Overall | 170 (129 HIV +/ 41 HIV -) | 443 | 241-776 | (-) 4.3 | (-)119.8, +111.2 | [5 empty] | 49 paired samples on Auto40 | median: 352 | 2.5% | Repeatability or intra-assay
[8 empty] | >200 | 33 | 83* | 38-137 | (-) 12.1 | (-) 62.1, +37.9 | [5 empty] | 49 paired samples on FACSCount | median: 362 | 2.9% | Repeatability or intra-assay
[8 empty] | 200-500 | 65 | 303* | 260-140 | (-) 7.4 | (-) 90.6, +75.9 | [5 empty] | 10 replicates of 3 blood samples | Cell counts: 23, 363, 572 | 7.1%, 1.1%, 2.4% | Intra-laboratory variation on FACSCount (inter-assay)_
[8 empty] | >500 | 72 | 841* | 678-1099 | 2.1 | (-)153.1, +157.3 | [5 empty] | 10 replicates of 3 blood samples | Cell Counts: 25, 355, and 557 | 13.3%, 1.7%, 2.0% | Intra-laboratory variation on Auto40 (inter-assay)
25 | Herbert 2013 | UK | Hospital Laboratory | One author associated with Alere, Alere provided test cartridges and machines | Beckman Coulter FC 500 | PIMA | HIV positive, newly diagnosed and follow-up patients | Total | 283 (only 254 paired results after lead-in period) | 450* | 10-1290 |  |  |  | 93.3%, 200 (n=30): PPV 75.7% (n=37) | 96.0%, 200 (n=224): NPV 99.1 (n=217)
[8 empty] | Newly Diagnosed | 47 | 435* | 60-1200 |  |  |  | 94.8%, 350 (n=96): PPV 82.7 (n=110) | 88.0%, 350 (n=158): NPV 96.5% (n=144)
[8 empty] | Chronic Infection | 236 | 445* | 10-1290 |  |  |  | 98.6%, 500 (n=142): PPV 80.9% (n=173) | 70.5%, 500 (n=112): NPV 97.5% (n=81)
26 | Logan 2013 | USA | Research laboratory | One author associated with Mbio, | FACSCalibur | Mbio SnapCount System | HIV positive | Venous blood, all | 94 | 541* | 19-1215 | (-) 10 | (-)132, +112 |  |  |  | 2.1% (n=2), 350 | 3.2% (n=3), 350 | 3 replicates, n= 85: intra-assay, venous blood | >350 | 7%
[8 empty] | venous blood <500 | 44 |  |  | (+)2 | (-)92, +96 | [5 empty] | 3 replicates, n=33: intra-assay, venous blood | <350 | 10%
[8 empty] | capillary blood, all | 52 | 541* | 19-1215 | (-)4 | (-)195, +186 |  |  |  | 3.8% (n=2) |  | 3 replicates, n=30, intra-assay, capillary blood | >350 | 13%
[8 empty] | capillary, <500 | 25 |  |  | 21 | (-)85, +126 | [5 empty] | 3 replicates, n=30, intra-assay, capillary blood | <350 | 11%
27 | Manabe 2012 | Uganda | Tertiary hostital laboratory | Study funded by Alere | BD FACSCalibur with TruCount beads | PIMA | HIV + | Overall | 380 | 396* | 18-1500
[8 empty] | Venous Blood, Overall | 206 |  |  | (-)68.5 | (-)79.6, (-)57.4 | 275 | 94.3%, <250 | 85.4%, <250
[8 empty] | 0-250 | 35 |  |  | 13.6 | 2.5, 24.7 |  | 98.2%, <300 | 75.3%, <300
[8 empty] | 250-500 | 107 |  |  | (-)54.6 | (-)64.5, (-)44.7
[8 empty] | >500 | 64 |  |  | (-)121.7 | (-)147.9, (-)95.4
[8 empty] | Capillary Blood, Overall | 176 |  |  | (-)66.3 | (-)83.4, (-)49.2 | 290 | 96.3%, <250 | 86.6%, <250
[8 empty] | 0-250 | 27 |  |  | (-)10.8 | (-)-27.3, +5.6 |  | 93.2%, <300 | 79.5%, <300
[8 empty] | 250-500 | 90 |  |  | (-)47.3 | (-)62.2, (-)32.5
[8 empty] | >500 | 59 |  |  | (-)120.6 | (-)162.8, (-)78.4
[8 empty] | Venous vs. Capillary |  |  |  | (-)3.8 | (-)17.75, 10.25
28 | Mbopi-Keou (A) 2012 | Cameroon | National Health Laboratory | None stated | FACSCalibur | PanLeucogating Auto40 Flow Cytometer | HIV positive and negative | Overall | 234 (adults + children) |  |  | (-)15 (from figure) | (-)275, 250 (from figure) | ~450 | 87%, 200c/ul | 98%, 200 c/ul
[8 empty] | Overall | 146 adults | 658 (SD 440) |  | 3.6 | (-)195.7, +202.9 |  | 92%, 350c/ul | 98%, 350c/ul
[8 empty] | Overall | 88 children | 1569 (SD 933) |  | 20.5 | (-) 318.5, 359.5 |  | 82%, 750 c/ul | 98%, 750c/ul
[8 empty] | <200 | 17 (adults + children) | 85 (62) |  | 18.8 | (-)66.8, 104.4 | [5 empty] | n=3 | 103, <200 | 8.7%, intrarun assay
[8 empty] | 200-350 | 34 (adults + children) | 283 (46) |  | 14 | (-)75.6, 103.6 | [5 empty] | n=3 | 415, 200-350 | 4.0%, intrarun assay
[8 empty] | >350 | 183 (adults + children) | 1219 (773) |  | 8.1 | (-)289.3, 305.5 | [5 empty] | n=3 | 921, >750 | 3.1%, intrarun assay
[19 empty] | n=3 | 122, <200 | 9.0%, interrun assay
[19 empty] | n=3 | 431, 20-350 | 4.8%, interrun assay
[19 empty] | n=3 | 937, >750 | 4.1%, interrun assay
29 | Mbopi- Keou (B) 2012 | Cameroon | HIV reference laboratory | None stated | FACSCalibur | Auto40 flowcytometer | Overall |  | 257 (adults + children) |  |  | 7.6 |  |  | 87%, 200 | 99%, 200 |  |  |  | 425 c/ul | 3.7%, intrarun
[15 empty] | 87%, 350 | 98%, 350 |  |  |  | 122 c/ul | 7.4% intrarun
[7 empty] | Volunteers in Yaounde |  | 188 (adults + children) | 989 (883) | (5-4454) | 6.8 | (-)12.4, 26.0 | 450 |  |  |  |  | 2 samples | mean of above two | 5.5% intra run
[7 empty] | Volunteers in Ambang Bikok |  | 69 (adults) | 1032 | 281-1616 | 9.7 | (-)19.9, 39.4 | 300 | [5 empty] | 439c/ul | 6.1%, interrun
[20 empty] | 133 c/ul | 9.7% interrun
[19 empty] | 2 samples | mean of above two | 7.9% interrun
Where HIV positive data can be extracted, this alone has been included. When this is not possible, combined HIVpos/neg data included.
* mean difference unless stated otherwise, + indicates index test gave a higher result, - indicates index test gave a lower result. M=median difference (where no + or – symbol is given, the figure represents the median absolute difference).
#Range of CD4 (read from plot, therefore approximate), mean, SD (all as measured by reference test, unless only BA plot available for estimating range, in which case as measured by mean of two tests) (cells/µl)
†read from plot, therefore approximate
‡Limits of agreement

## Data Overview
| Bias reported? | Bias of subgroup | BA | Misclass | Within-lab precision whole blood (absolute CD4 count) | Within-lab precision stabilised blood
Balakrishnan 2004 | 1 | 1 | 1 | 1 | 0
Carella 1995 | 0 | 0 | 0 | 1 | 0
Landay 1993 | 0 | 0 | 0 | 1 | 1
Gernow 1995 | 0 | 0 | 0 | 1 | 1
Didier 2001 | 1 | 0 | 0 | 0 | 0
Karcher 2006 | 1 | 1 | 1 | 1 | 0
Lutwama 2008 | 1 | 1 | 1 | 1 | 0
Diabouga 2003 | 1 | 1 | 0 | 0 | 1
Imade 2005 | 0 | 0 | 0 | 0 | 1
Idigbe 2006 | 0 | 0 | 0 | 1 | 0
Balakrishnan 2006 | 1 | 1 | 1 | 1 | 0
Kandathil 2005 | 1 | 0 | 1 | 1 | 1
Spacek 2006 | 1 | 1 | 1 | 1 | 0
Thakar 2006 | 1 | 0 | 1 | 1 | 1
Pattanapanyasat 2007 | 1 | 1 | 1 | 0 | 0 | 1
Pattanapanyasat 2008 | 1 | 1 | 1 | 0 | 0 | 1
Zijenah 2006 | 1 | 1 | 1 | 1 | 0
Fryland 2006 | 1 | 0 | 1 | 0 | 1
Cassens 2004 | 1 | 0 | 1 | 0 | 0
Pattanapanyasat 2005 | 1 | 1 | 1 | 0 | 0 | 1
Manasa 2007 | 1 | 0 | 1 | 0 | 1
Dieye 2005 | 1 | 0 | 1 | 0 | 0
Lynen 2006 | 1 | 1 | 1 | 1 | 1
Strauss 1996 | 0 | 0 | 0 | 0 | 1
Lopez 1999 | 0 | 0 | 0 | 0 | 1
Johnson 1995 | 0 | 0 | 0 | 0 | 1
Nicholson 1994 | 0 | 0 | 0 | 0 | 1
Schnizlein-Bick 2000 | 1 | 1 | 0 | 0 | 1
Nicholson 1997 | 1 | 0 | 0 | 0 | 1
Higgins 2007 | 1 | 0 | 1 | 0 | 0 |  | Higgins reported %CV for CD4% rather than absolute CD4
Jeganathan 2008 | 1 | 0 | 1 | 0 | 0
Reimann 2000 | 1 | 0 | 0 | 0 | 1
Storie 2004 | 1 | 1 | 1 | 0 | 0
Storie 2003 | 1 | 1 | 1 | 0 | 0
Denny 2008 | 1 | 1 | 0 | 0 | 1
Pattanapanyasat, 2005 | 1 | 1 | 1 | 0 | 0 | 1
Glencross, 2001 | 1 | 0 | 1 | 0 | 0
Sippy-Chatrani, 2008 | 1 | 1 | 1 | 0 | 0
Briggs 2007 | 1 | 0 | 1 | 1 | 1
Srithanaviboonchai 2008 | 1 | 0 | 1 | 0 | 0
Lyamuya 1996 | 0 | 0 | 0 | 0 | 0
Mtapuri-Zinyowera 2010 | 1 | 1 | 1 | 1 | 0
Jani 2011 | 1 | 1 | 1 | 1 | 1
Sukapirom 2011 | 1 | 1 | 1 | 0 | 1
Renault 2010 | 1 | 1 | 0 | 1 | 0
Bergeron 2010 | 1 | 1 | 1 | 0 | 0 | 1
Sitoe 2011 | 1 | 1 | 1 | 0 | 1
| 37 | 23 | 30 | 17 | 21
total | 77.0833 | 47.9167 | 62.5 | 35.4167 | 43.75
=repeat appearances of same study (ie study evaluates more than one technology)