SUPPLEMENTAL INFORMATION

Table S1. List of potential interferences for analyzing 52Cr in urine.  Masses determined using ThermoFinnigan Element2 ICP-MS Interference Workshop software.

| 52Cr (m/z) | Potential Interference (AB) | Isotopic Abundance A % | Expected Concentration of A in Urine (µg L-1) | Isotopic Abundance B % | Expected Concentration of B in Urine (µg L-1) |
| --- | --- | --- | --- | --- | --- |
| 51.4528 | 103Rh++ | 100 | 0.001 - 0.039 | - | - |
| 51.9520 | 104Pd++ | 11.1 | 0.07 - 0.64 | - | - |
| 51.9527 | 104Ru++ | 18.7 | Expected to be low | - | - |
| 51.9620 | 36S16O+ | 0.02 | High | 99.8 | atmosphere |
| 51.9624 | 12C40Ar+ | 98.9 | High | 99.6 | ICP gas |
| 51.9625 | 36Ar16O+ | 0.3 | ICP gas | 99.8 | atmosphere |
| 51.9658 | 14N38Ar+ | 99.6 | atmosphere | 0.06 | ICP gas |
| 51.9680 | 35Cl17O+ | 75.8 | ≤ 5,900,000 | 0.04 | atmosphere |
| 52.4530 | 105Pd++ | 22.3 | 0.07 - 0.64 | - | - |
| 51.9518 | 51V1H+ | 99.8 | 0.2 - 10 | 100 | atmosphere |
| 51.9626 | 40Ca12C+ | 96.9 | 30,000 – 200,000 | 98.9 | ~ 5,000,000 |
| 51.9671 | 39K13C+ | 93.3 | 542,000 - 9,770,000 | 1.1 | ~ 5,000,000 |
| 51.9660 | 37Cl15N+ | 24.2 | ≤ 5,900,000 | 0.4 | atmosphere |
| 51.9768 | 37Cl14N1H+ | 24.2 | ≤ 5,900,000 | 99.6/100 | atmosphere |
| 51.9716 | 35Cl16O1H+ | 75.8 | ≤ 5,900,000 | 99.8/100 | atmosphere |
| 51.9699 | 33S19F+ | 0.8 | High | 100 | 200 - 3,200 |
| 51.9676 | 21Ne31P+ | 0.3 | Expected to be low | 100 | 220,000 - 2,600,000 |
| 51.9645 | 20Ne32S+ | 90.5 | Expected to be low | 95.0 | ~ 5,000,000 |

Table S2. List of potential interferences for analyzing 60Ni in urine. Masses determined using ThermoFinnigan Element2 ICP-MS Interference Workshop software.

| 60Ni (m/z) | Potential Interference (AB) | Isotopic Abundance A % | Expected Concentration of A in Urine (µg L-1) | Isotopic Abundance B % | Expected Concentration of B in Urine (µg L-1) |
| --- | --- | --- | --- | --- | --- |
| 59.4516 | 119Sn++ | 8.6 | 0.05 - 2.28 | - | - |
| 59.9504 | 44Ca16O+ | 2.1 | 30,000 - 200,000 | 99.8 | atmosphere |
| 59.9511 | 120Sn++ | 32.6 | 0.05 - 2.28 | - | - |
| 59.9520 | 120Te++ | 0.1 | Expected to be low | - | - |
| 59.9526 | 24Mg36Ar+ | 79 | 40,500 - 243,000 | 0.3 | ICP gas |
| 59.9541 | 22Ne38Ar+ | 9.2 | Expected to be low | 0.06 | ICP gas |
| 59.9548 | 20Ne40Ar+ | 90.5 | Expected to be low | 99.6 | ICP gas |
| 60.4519 | 121Sb++ | 57.3 | 0.12 - 10.0 | - | - |
| 59.9410 | 59Co1H+ | 100 | 0.28 - 2.0 | 100 | atmosphere |
| 59.9495 | 58Ni2H+ | 68.3 | 0.055 - 20.0 | 0.02 | atmosphere |
| 59.9480 | 48Ti12C+ | 73.8 | 0.1 - 3.7 | 98.9 | ~ 5,000,000 |
| 59.9525 | 48Ca12C+ | 0.2 | 30,000 - 200,000 | 98.9 | ~ 5,000,000 |
| 59.9551 | 47Ti13C+ | 7.3 | 0.1 - 3.7 | 1.1 | ~ 5,000,000 |
| 59.9557 | 46Ti14N+ | 8.0 | 0.1 - 3.7 | 99.6 | atmosphere |
| 59.9560 | 45Sc15N+ | 100 | - | 0.4 | atmosphere |
| 59.9668 | 45Sc14N1H+ | 100 | - | 99.6/100 | atmosphere |
| 59.9615 | 43Ca16O1H+ | 0.1 | 30,000 - 200,000 | 99.8/100 | atmosphere |
| 59.9688 | 41K18O1H+ | 6.7 | 542,000 - 9,770,00 | 0.2/99.9 | atmosphere |
| 59.9602 | 41K19F+ | 6.7 | 542,000 - 9,770,00 | 100 | 200 - 3,200 |
| 59.9503 | 29Si31P+ | 4.7 | 100 - 51,600 | 100 | 220,000 - 2,600,000 |
| 59.9490 | 28Si32S+ | 92.2 | 100 - 51,600 | 95.0 | ~ 5,000,000 |
| 59.9505 | 26Mg34S+ | 11.0 | 40,500 - 243,000 | 4.2 | ~ 5,000,000 |
| 59.9521 | 24Mg36S+ | 79.0 | 40,500 - 243,000 | 0.02 | ~ 5,000,000 |
| 59.9530 | 27Al33S+ | 100 | 5 - 30 | 0.8 | ~ 5,000,000 |
| 59.9547 | 25Mg35Cl+ | 10.0 | 40,500 - 243,000 | 75.8 | ~ 6,000,000 |
| 59.9557 | 23Na37Cl+ | 100 | 281 - 10,110 | 24.2 | ~ 6,000,000 |

Figure S1a.  The effects of cell gas flow rate on ion signal intensity at m/z 52 using methane gas on Cr analysis in urine samples spiked with 1 µg L-1 Cr, 1% ethanol (40Ar12C+), and/or 1% HCl (35Cl16O1H+).

Figure S1b.  The effects of cell gas flow rate on ion signal intensity at m/z 52 using oxygen gas on Cr analysis in urine samples spiked with 1 µg L-1 Cr, 1% ethanol (40Ar12C+), and/or 1% HCl (35Cl16O1H+).

Figure S1c.  The effects of cell gas flow rate on ion signal intensity at m/z 52 using ammonia gas on Cr analysis in urine samples spiked with 1 µg L-1 Cr, 1% ethanol (40Ar12C+), and/or 1% HCl (35Cl16O1H+).

Figure S1d.  The effects of cell gas flow rate on ion signal intensity at m/z 52 using argon/hydrogen gas on Cr analysis in urine samples spiked with 1 µg L-1 Cr, 1% ethanol (40Ar12C+), and/or 1% HCl (35Cl16O1H+).

Figure S1e.  The effects of cell gas flow rate on ion signal intensity at m/z 52 using helium gas on Cr analysis in urine samples spiked with 1 µg L-1 Cr, 1% ethanol (40Ar12C+), and/or 1% HCl (35Cl16O1H+).

Figure S2a.  The effects of cell gas flow rate on ion signal intensity at m/z 60 using methane gas on Ni analysis in urine samples spiked with 1 µg L-1 Ni and/or 1% ethanol + 20 mg L-1 Ca (44Ca16O+).

Figure S2b.  The effects of cell gas flow rate on ion signal intensity at m/z 60 using oxygen gas on Ni analysis in urine samples spiked with 1 µg L-1 Ni and/or 1% ethanol + 20 mg L-1 Ca (44Ca16O+).

Figure S2c.  The effects of cell gas flow rate on ion signal intensity at m/z 60 using ammonia gas on Ni analysis in urine samples spiked with 1 µg L-1 Ni and/or 1% ethanol + 20 mg L-1 Ca (44Ca16O+).

Figure S2d.  The effects of cell gas flow rate on ion signal intensity at m/z 60 using argon/hydrogen gas on Ni analysis in urine samples spiked with 1 µg L-1 Ni and/or 1% ethanol + 20 mg L-1 Ca (44Ca16O+).

Figure S2e.  The effects of cell gas flow rate on ion signal intensity at m/z 60 using helium gas on Ni analysis in urine samples spiked with 1 µg L-1 Ni and/or 1% ethanol + 20 mg L-1 Ca (44Ca16O+).

Figure S3.  Displays the response at m/z 52 for 2% HNO3, 1% EtOH, and/or 1 µg L-1 Cr under increasing helium cell gas flow rates in KED mode.

Figure S4.  Displays the response at m/z 60 for 2% HNO3, 1% EtOH and 20 mg L-1 Ca, and/or 1 µg L-1 Ni under increasing helium cell gas flow rates in KED mode.

Table S3. ESI FAST method control parameters that control the functions for the autosampler events during the NexION ICP-MS wash time.

| Event | Action | Parameters | Parameter Units |
| --- | --- | --- | --- |
| On Probe Down | Vacuum1 On |  |  |
| On Probe Down | Load1 |  |  |
| Probe In Sample | Timer A | 4 | seconds |
| Timer A Expires | Inject1 |  |  |
| Timer A Expires | Move Rinse |  |  |
| Rinse Completed | Probe Up |  |  |
| On Rinse | Vacuum1 On |  |  |
| On Rinse | Load1 |  |  |
| On Rinse | Probe Down |  |  |
| On Rinse | Timer B | 2 | seconds |
| Timer B Expires | Probe Up |  |  |
| Timer B Expires | Timer C | 2 | seconds |
| Timer C Expires | Probe Down |  |  |
| Timer C Expires | Timer D | 2 | seconds |
| Timer D Expires | Probe Up |  |  |
| Timer D Expires | Timer E | 2 | seconds |
| Timer E Expires | Probe Down |  |  |
| Timer E Expires | Timer F | 2 | seconds |
| Timer F Expires | Move Next |  |  |

Table S4.  In house CDC quality control samples used to show % recovery of the method.  N = 5.

|  |  | Cr |  |  |
| --- | --- | --- | --- | --- |
| Sample ID |  | Target (µg L-1) | Average ± 1 SD (µg L-1) | % Recovery ± 1 SD |
| Low QC |  | 0.820 | 0.885 ± 0.128 |  |
| + 1 µg L-1 |  | 1.82 | 1.75 ± 0.13 | 96.3 ± 7.1 |
| + 3 µg L-1 |  | 3.82 | 3.87 ± 0.53 | 101 ± 14 |
| + 10 µg L-1 |  | 10.8 | 10.6 ± 0.2 | 98.4 ± 1.9 |
|  |  | Ni |  |  |
| Sample ID |  | Target (µg L-1) | Average ± 1 SD (µg L-1) | % Recovery ± 1 SD |
| Low QC |  | 1.36 | 1.59 ± 0.10 |  |
| + 1 µg L-1 |  | 2.36 | 2.44 ± 0.12 | 103 ± 5 |
| + 3 µg L-1 |  | 4.36 | 4.38 ± 0.40 | 101 ± 9 |
| + 10 µg L-1 |  | 11.4 | 11.1 ± 0.3 | 97.9 ± 2.6 |